Purpose: Document the methodology used for correcting h2o and LE data measured at TP39 and TP02 during the middle of 2010.
Background:
- Li-7000s at TP02 and TP39 were calibrated in detail on March 20, 2010
- During this calibration, the IRGAs were improperly calibrated for h2o span.
- The error likely occurred as low-h2o air was passed through the IRGA when it was spanned for h2o, instead of reasonably-high h2o air.
- This resulted in a large increase in the h2o span calibration coefficient (S) in both IRGAs. TP39 IRGA Sh2o value increased from 1 to 1.4, and for TP02 from 1.02 to 1.36.
- This resulted in an increased gain in measured h2o values throughout the growing season, compared to what should be expected. Commonly, the h2o concentrations at both sites maxed out (originally 40 mmol/mol, then upped to 60 mmol/mol), leaving a good deal of half hours with concentrations fixed to this maximum value.
- This increased gain also caused an increase in LE values relative to other years (since any change in h2o was magnified, thus magnifying LE).
- The IRGAs were re-calibrated on October 25, 2010.
Correction Methodology:
(NOTE: It may be helpful to follow along in the MATLAB script: 'correcting_2010_LE_data.m', where all calculations and figures are done):
1. The Li-Cor Comp7000.exe program was used to recreate the conditions in the Li-7000 with both correct and faulty span values. A sequence of absorbance (Raw W) values were used to predict H20 measured by the Li-7000 at both span settings. These tables can be found in
   /1/fielddata/Documentation/Jay_IRGA_diagnostics_TP02_H2O_errors.ods and
   /1/fielddata/Documentation/Jay_IRGA_diagnostics_TP39_H2O_errors.ods
2. Figures 1 and 2 show a linear relationship between the 'correct' and 'faulty' values of h2o obtained using the same absorbance value. Data from both IRGAs show a linear relationship, with a non-zero intercept, implying that a simple ratio adjustment would not work for h2o.
   The following shows the slopes and intercepts between the 'correct' and 'faulty' span h2o values. Both slopes are >1 and both intercepts are < 0, indicating that the 'faulty' h2o values overestimate the real h2o values at medium and high h2o levels, but underestimate at low h2o levels.
  TP39: slope = 1.7807 and intercept = -1.8670
  TP02: slope = 1.4544 and intercept = -4.2219
3. To see the effect that these offsets have on calculated LE, What we've done is insert lines of code into fr_calc_eddy.m (the eddy flux calculation script that runs when we click “Recalc Fluxes from High Freq Data” in the data GUI), that will correct the h2o values (by linear relationship) and recalculate the LE fluxes. Note that these lines have been commented back out of the fr_calc_eddy file. The old and new values were outputted to the following files:
   /1/fielddata/Matlab/Data/Other/Fixing_2011_Li7000_cal/TP02_LE_test.txt
   /1/fielddata/Matlab/Data/Other/Fixing_2011_Li7000_cal/TP39_LE_test.txt
4. Figures 3 and 4 show the relationship between 'faulty' and 'correct' calculated LE values for both TP39 and TP02. Fortunately, the relationships are linear, with zero slopes. This means that we can correct the 'faulty' LE data by multiplying it by the slope between 'correct' and 'faulty' data:
   TP39 correction factor: 0.57122
   TP02 correction factor: 0.68951
5. Note that we can only correct the h2o and LE data during periods when the h2o values are NOT maxed out; otherwise, the measured data is meaningless, and we have to discard it.